Device name

CeraFix Bone Cement

- Contraindication

(1) Active infection;

(2) Hemorrhagic diseases;

(3) Patients with symptoms of spinal cord compression caused by pathological tissue;

(4) Individuals who are allergic to the ingredients in the product.

- Warning and Precaution

- Precaution:

1.The implantation process and conditions of the product should comply with the relevant technical management regulations of the National Health Commission (NHC).

2. Surgical surgeons must receive training on the operating techniques of this product and carefully follow them. Operators must understand and follow the recommended time required for each stage of preparation. If the operating instructions are not followed, adverse reactions may occur.

3.The injection of bone cement must be slowly carried out under the control of continuous X-ray monitoring, observing the distribution at the lesion site and paying attention to any leakage outside the lesion site. In case of bone cement leakage, the operation should be stopped immediately.

4.Before the operation, the patient must undergo a thorough preoperative examination.

5. Please ensure the normal operation of the ventilation system in the operating room and try to eliminate the gas emitted by monomer fluids as much as possible. Monomer liquid is a volatile and flammable liquid.

6.If other objects are implanted near the surgical site after surgery, it may increase the risk of infection.

7.According to the Expert Consensus on the Diagnosis and Treatment of Osteoporotic Vertebral Compression Fractures (2021 Edition), the integrity of the fractured vertebral peripheral wall needs to be assessed: preoperative assessment of the integrity of the vertebral peripheral wall by CT reconstruction is particularly important, and appropriate countermeasures should be formulated preoperatively in cases of peripheral wall disruption.

8.For patients with peripheral wall breakage, the operator is required to have rich surgical experience, and the application of temperature gradient perfusion technique, bone cement split modulation, sealing and bone anchoring technique can improve the safety of the operation and reduce the complication rate.

Warning:

1.The length of time required for each stage of operation is not only related to the level of room temperature and the temperature of each material, but also to the humidity of the operating room. At high temperatures, the hardening time will be reduced, while at low temperatures, the required time will be extended.

2. Please carefully read the IFU before use.

3. Operators should receive specialized training and relevant experience to thoroughly familiarize themselves with the properties of the product, the characteristics of operation, the applicability of the product, and the implantation method of bone cement.

4. Our company does not recommend surgical techniques, and it is up to the operator to decide how to effectively use this product and the operating techniques for patients.

5. Strictly follow the instructions for the use, mixing, and preparation of bone cement.

6. It is strictly prohibited to sterilize the product again, and it is only for one-time use. Only when the original packaging is unopened and undamaged can complete sterilization be ensured.

7.If there is a hypotension reaction between 10 and 165 seconds. Some patients may even worsen to cardiac arrest. Therefore, it is important to carefully monitor changes in blood pressure during or after the placement of bone cement.

8. Methyl methacrylate may cause allergic reactions in certain high-risk groups of patients.

9.Due to insufficient clinical data, it is uncertain whether this product is harmless for pregnant, lactating women, and children.

10.The established principles and techniques of surgical procedures must be strictly adhered to. Deep wound infection is a serious postoperative complication and may necessitate the complete removal of the implanted bone cement.

11. It is necessary to strictly prevent excessive exposure to harmful gases from monomer solutions, which may irritate the respiratory tract, eyes, and even the liver.

12. Before starting the operation, always check the status of the packaging contents. Do not use if there is turbidity or premature aggregation of the liquid, or if the powder color has turned yellow.

13. Do not let monomer solution come into contact with rubber or latex gloves. This solution is a powerful solvent, and in case of contact, the gloves may be dissolved, causing tissue damage. Wearing two pairs of gloves may reduce the likelihood of allergic reactions.

14. Do not allow personnel wearing contact lenses to approach or be responsible for mixing bone cement.

15. Use appropriate imaging techniques to confirm the correct position of the puncture needle, the absence of damage to surrounding tissues, and the accurate placement of injected bone cement. Use imaging techniques such as fluoroscopy to measure the volume of bone cement filled into the spine.

16. Avoid excessive compression of bone cement to prevent leakage to other areas that are not intended for treatment. The extravasation of bone cement can cause damage to surrounding tissues and problems with the nervous and circulatory systems.

17.During the injection period, if the puncture needle is inserted into the vein or there are undetected micro cracks, it can cause leakage.

18.During the operation, if any leakage of bone cement is found outside the spine or in the circulatory system, filling should be stopped immediately.

19. The final polymerization reaction stage occurs when filling the original site, which generates an exothermic reaction and releases a considerable amount of heat. Therefore, it is necessary to maintain the patient's position until the polymerization process is completed, so that the bone cement can be properly fixed. This may require an additional 1 to 2 hours or more, depending on the patient's health condition, and it is up to the operator to decide.

20. Monomer liquids have high volatility and flammability, therefore appropriate preventive measures should be taken in the operating room to maintain good ventilation.

21. The pouch of the powder container is additionally packaged with non sterile protective aluminum foil.