Device name

Bone Filling Container and Bone Filling Container Kit

Contraindication

1) Osteomyelitis, epidural cyst;
2) Vertebral compression, more than 70%
3) Vertebra spinal cord severely depressed or secondary spinal steno-sis;
4) Severe cardiopulmonary dysfunction, blood coagulation disorders, sepsis and other surgery is not suitable. 
 Warnings and precautions 

Bone cement is generally in accordance with a bag of powder (10 g): 1 bottle of liquid (5ml) ratio deployment to constantly stir it evenly 30s ~ 60s after a while, and then poured into a syringe injection, injection period (working period) about 1 to 4 minutes after the deployment, see the bone cement manual.

Pay attention to the packages of the instruments before use. If the damage of packages is found, do not use.
The valid period of sterilization of the product is 24 months, please use within the validity period.
When the completion of the injection, immediately rotate counterclockwise the handle 1 turn to release residual pressure, preventing accidental injection of bone cement.
Do not reuse, reusing may cause cross contamination.
The parts of Bone Filling Container after clinical used shall be disposed in accordance with the relevant medical waste management regulations

This device is a permanent implant. The avoidance of the usage of patient with highly allergy to metal or artificial vessel fabric should be noticed.

This product has been evaluated and can be used in a 3.0T magnetic resonance (MRI) environment.

This instruction manual must be read carefully before use;

The doctor evaluates the patient before surgery, and determines whether the patient is suitable for the product according to his health condition, mental condition, allergy to foreign bodies, etc. if applicable, choose the applicable models and specifications according to the patient's own situation, and different models and specifications of products should be prepared before the operation.

During the process of use, the doctor should wear medical gloves for operation.

In the event of a serious incident, immediate medical treatment must be provided to the injured. The faulty device, usage records, and key information should be preserved. The user or medical staff must report the incident to the manufacturer and the competent authority of the member state where the user and patient are located.

